Technical specifications

AGATA GE DETECTOR

| Ind. | Date | Historic of modifications |
| --- | --- | --- |
| A | 10/12/2021 | Original issue |
|  | 04/02/2022 | update |

Table of contents

1. General description	3

2.	The detector:	3

2.1	Crystals	3

2.2	Crystal specification	3

2.3	Crystallographic axes	3

2.4	Crystal size	3

2.5  Crystal shape:	3

2.6  Crystal segmentation:	4

3.	The capsule	4

3.1. Capsule:	4

3.2. Wall thickness	4

3.3. Drawing:	4

3.4. Segment cabling:	4

3.5. Material:	4

4.	Detector polarization:	4

5.	Timing	4

5.1. Delivery:	4

5.2.	Acceptance:	5

5.3.	Warranty:	5

6.	Factory acceptance tests:	5

6.1.	Full volume:	5

6.2.	For the segments:	5

6.3	Efficiency:	5

7.	Customer acceptance tests:	5

7.1	Full volume:	6

7.2  For All the segments:	6

7.3	Efficiency:	6

8.	Annealing of the detector	6

# 1. General description

AGATA is a detection unit made of three HPGe detectors in semi-coaxial configuration sharing a common cryostat. Each crystal has a pyramid shape: the front-end features a section shaped like an irregular hexagon but the opposite (rear) end has still a cylindrical shape. The crystals have three slightly different shapes. Their external surface is electrically segmented into 36 parts. Each crystal is mounted in a vacuum-tight capsule.

# The detector:

## Crystals

The crystals will be made of n-type HPGe, with an impurity concentration between 0.3 and 1.8 1010/cm3. The maximum impurity concentration has to be placed at the front end of the capsule.  The manufacturer will provide a measurement of impurities concentration (accuracy of +/- 0.05 1010/cm3) at the centre and sides of the two ends (front and rear).

## Crystal specification

Impurity concentration on front and back parts as well as capacity and resistivity as a function of the applied high voltage will be submitted by the manufacturer.

## Crystallographic axes

The 100 Germanium crystallographic axis will pass through the central corner of the three capsules when assembled as a triplet with an accuracy of +/- 3°. Crystal orientation will be marked on the external capsule's front end or be provided on specific drawings. Crystal orientation will be the same for all crystals.

## Crystal size

If one dimension is out of the specifications, the manufacturer will inform the AGATA Management Board. An agreement will be discussed if the efficiency of the detector is acceptable.

The crystal sizes will be as follows:

- length 90 (+0.6 -1.4) mm

- rear side diameter 80 (+0.6 -0.6) mm

Crystal central hole: nominal diameter 10 mm extends to 13 mm from the front end.

The net Germanium weight will be provided from the manufacturer.

## 2.5 Crystal shape:

Crystal will have 6 side faces that are flat and narrowing, forming an irregular hexagon on the detector's front end. The cylindrical part at the rear end must be at least 5 mm long.

There are three different crystal shapes (named A/red, B/green and C/blue) as defined in the AGATA collaboration publication in Figure 1 and 2 on p. 224 and 225 (Nuclear Instruments and Methods in Physics Research A 618 (2010) 223–233).

## 2.6 Crystal segmentation:

The crystals' external electrode will be segmented into 36 parts with 6 longitudinal segmentations and 5 transversal segmentations. Longitudinal segmentation lines will be on the middle of the narrowing external faces. Transversal segmentation lines will divide the detector into parallel sections with a width of 8, 13, 15, 18, 18 and 18mm, respectively, starting with 8mm from the hexagonal end of the crystal. The outer contact segmentation gap width will be a nominal of 0.5 mm.

# The capsule

## 3.1. Capsule:

Each crystal will be mounted in a vacuum tight capsule.

A unlimitedly re-usable canister model is requested.

## 3.2. Wall thickness

The distance between capsule walls and crystal side faces will be 0.4 - 0.7 mm. The thickness of the front and walls of the capsule must be 0.8 +/- 0.05 mm.

## 3.3. Drawing:

The manufacturer will submit a drawing of the capsule's rear end showing the position of vacuum feedthroughs and threaded holes for detector and board mounting.

## 3.4. Segment cabling:

The manufacturer will supply the drawing of cabling of the segments at the vacuum feedthroughs. The cabling scheme will be the same for all capsules. The manufacturer will provide the capacitance and inductance of the cabling used inside the capsules.

## 3.5. Material:

There will be no material containing fluorine inside the capsule.

# Detector polarization:

The operational voltage Vop will be at least the following:

Vop ( (10 x Vdepl + 7500) / 11

where:

Vop	=	operating voltage

Vdepl	=	depletion voltage

V in volts

Maximum operating voltage will be 5000 V.

# Timing

## 5.1. Delivery:

Within 18 months after the signature of the ordering.

## Acceptance:

The AGATA detector acceptance test will consist of a factory acceptance test (FAT) and eventually a customer acceptance test (CAT) if AGATA collaboration needs it. In this case the CAT has to be finished within two months after delivery. The FAT and CAT are described in section 6 and 7.

## Warranty:

The manufacturer has to provide a full warranty for a minimum of two years starting with the FAT, and CAT if present.

# Factory acceptance tests:

These will be performed using a test cryostat equipped with 37 cold input stages and using electronics from the manufacturer.

## Full volume:

Measurement of the FWHM for 1.33MeV (60Co):

Guaranteed value: ( 2.35keV

Measurement of the FWHM at 122keV (57Co):

Guaranteed value: ≤ 1.35keV

Peak shape: FWTM / FWHM:

Guaranteed value: ( 2.0

Measurement conditions:

The source is positioned in front of the detector,

Count rate ( 1000 counts per second,

Gaussian shaping time: 6µs.

## For the segments:

Measurement of the FWHM at 1.33MeV (60Co):

Guaranteed value:  ( 2.30keV

Measurement of the FWHM at 60keV(241Am):

Guaranteed value: ( 1.30keV

Mean value: ( 1.20keV

All segments will be measured at 60keV and at 1.33MeV

Measurement conditions:

The source is in front of each segment.

Count rate ( 1000 counts per second.

Gaussian shaping time: 6µs.

## Efficiency:

The efficiency of the detector at 1.33 MeV will be measured with a 60Co source at 25 cm distance from the Ge front end.

# Customer acceptance tests:

These will be performed using a test cryostat of the AGATA Collaboration equipped with 37 cold input stages and using commercial electronics.

Any non-compliance observed during CAT has to be justified by the customer as not related to a defect test cryostat. These tests will be carried out within 2 months of the detector being received.

To map out the Customer Acceptance Tests, the serial number of the test cryostat and electronics used will be provided by the customer to the manufacturer as well as operating conditions.

The following information will be provided by the customer to the manufacturer.

## Full volume:

Measurement of the FWHM for 1.33MeV (60Co):

Guaranteed value: ( 2.35keV

Measurement of the FWHM at 60keV (241Am):

Guaranteed value: ( 1.30keV

Peak shape: FWTM / FWHM:

Guaranteed value: ( 2.0

Measurement conditions:

The source is positioned in front of the detector.

Count rate ( 1000 counts per second

Gaussian shaping time: 6µs.

## 7.2  For All the segments:

Measurement of the FWHM at 1.33MeV (60Co):

Guaranteed value:  ( 2.30keV

Mean value:  ( 2.15 keV

Measurement of the FWHM at 60keV(241Am):

Guaranteed value: ( 1.30 keV

Mean value: ( 1.20keV

Measurement conditions:

The source is in front of each segment.

Count rate ( 1000 counts per second.

Gaussian shaping time: 6µs.

## Efficiency:

The efficiency of the detector at 1.33 MeV will be measured with a 60Co source at a distance of 25 cm from the front of the Ge crystal.

# Annealing of the detector

Before delivery, the supplier will perform an annealing cycle of the encapsulated detector which includes at least 10 hours at 100°C. The detector is guaranteed to withstand further annealing cycles at the customer up to a maximum temperature of 102 0C within the warranty period. Any defect occurring during annealing, as long as annealing is being performed under following conditions at customer’s site, will be covered by warranty period.

Annealing procedure:  Maximum temperature increase: 10 degrees per hour

Maximum temperature: 102 0C